T14. CirCor Clinical Outcome Results

CirCor DigiScope 2022 clinical outcome (normal versus abnormal), patient-grouped 5-fold cross-validation with a nested per-fold search, scored at recording level and at patient level under each aggregation rule. Mean and SD over the 5 folds.

| Model | Level | Rule | Folds | Units / fold | Sensitivity | Sensitivity SD | Specificity | Specificity SD | Balanced accuracy | Balanced accuracy SD | F1 | ROC-AUC | Accuracy |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| M3 | recording | none | 5 | 632.6 | 0.446 | 0.067 | 0.758 | 0.037 | 0.602 | 0.031 | 0.521 | 0.623 | 0.607 |
| M4 | recording | none | 5 | 632.6 | 0.483 | 0.068 | 0.718 | 0.031 | 0.600 | 0.024 | 0.540 | 0.628 | 0.604 |
| M5 | recording | none | 5 | 632.6 | 0.491 | 0.071 | 0.700 | 0.032 | 0.595 | 0.029 | 0.540 | 0.627 | 0.599 |
| M6 | recording | none | 5 | 632.6 | 0.457 | 0.104 | 0.736 | 0.073 | 0.597 | 0.023 | 0.521 | 0.632 | 0.601 |
| M7 | recording | none | 5 | 632.6 | 0.457 | 0.104 | 0.736 | 0.073 | 0.597 | 0.023 | 0.521 | 0.632 | 0.601 |
| M3 | patient | max | 5 | 188.4 | 0.671 | 0.066 | 0.537 | 0.048 | 0.604 | 0.026 | 0.619 | n/a | 0.602 |
| M4 | patient | max | 5 | 188.4 | 0.726 | 0.056 | 0.463 | 0.051 | 0.595 | 0.013 | 0.631 | n/a | 0.590 |
| M5 | patient | max | 5 | 188.4 | 0.766 | 0.062 | 0.393 | 0.013 | 0.579 | 0.026 | 0.634 | n/a | 0.573 |
| M6 | patient | max | 5 | 188.4 | 0.673 | 0.054 | 0.537 | 0.054 | 0.605 | 0.023 | 0.621 | n/a | 0.603 |
| M7 | patient | max | 5 | 188.4 | 0.673 | 0.054 | 0.537 | 0.054 | 0.605 | 0.023 | 0.621 | n/a | 0.603 |
| M3 | patient | mean | 5 | 188.4 | 0.480 | 0.084 | 0.747 | 0.042 | 0.614 | 0.045 | 0.547 | n/a | 0.618 |
| M4 | patient | mean | 5 | 188.4 | 0.531 | 0.074 | 0.706 | 0.027 | 0.618 | 0.035 | 0.574 | n/a | 0.621 |
| M5 | patient | mean | 5 | 188.4 | 0.524 | 0.075 | 0.701 | 0.045 | 0.613 | 0.039 | 0.567 | n/a | 0.616 |
| M6 | patient | mean | 5 | 188.4 | 0.485 | 0.074 | 0.753 | 0.022 | 0.619 | 0.037 | 0.553 | n/a | 0.623 |
| M7 | patient | mean | 5 | 188.4 | 0.485 | 0.074 | 0.753 | 0.022 | 0.619 | 0.037 | 0.553 | n/a | 0.623 |
| M3 | patient | any_present | 5 | 188.4 | 0.671 | 0.066 | 0.537 | 0.048 | 0.604 | 0.026 | 0.619 | n/a | 0.602 |
| M4 | patient | any_present | 5 | 188.4 | 0.726 | 0.056 | 0.463 | 0.051 | 0.595 | 0.013 | 0.631 | n/a | 0.590 |
| M5 | patient | any_present | 5 | 188.4 | 0.766 | 0.062 | 0.393 | 0.013 | 0.579 | 0.026 | 0.634 | n/a | 0.573 |
| M6 | patient | any_present | 5 | 188.4 | 0.695 | 0.093 | 0.490 | 0.109 | 0.592 | 0.013 | 0.619 | n/a | 0.589 |
| M7 | patient | any_present | 5 | 188.4 | 0.695 | 0.093 | 0.490 | 0.109 | 0.592 | 0.013 | 0.619 | n/a | 0.589 |

The outcome label is NOT in training_data.csv: it is parsed from the '#Outcome:' line of each patient's .txt file.

Abnormal is the positive class and makes up 48.4% of the patients in an average test fold -- close to balanced, the only near-balanced task in this project -- but sensitivity and balanced accuracy still lead.

Models: M3, M4, M5, M6, M7. The shipped binary model M1 is not among them: the CirCor tracks carry their own model set.

CirCor labels a PATIENT and the model scores a RECORDING, so every recording inherits its patient's label, including recordings from locations where the finding is not audible. Recording-level figures are pessimistic by an unmeasured amount; see circor_label_propagation.md.

Patient-level rows aggregate a patient's recordings under three rules (max, mean, any_present). Aggregation moves the operating point; no rule is declared the winner, because that is a clinical judgement.

CirCor is a predominantly paediatric cohort; nothing here transfers a PhysioNet result to it (that is EXP-D1, T16).

PV-MEPCG / PulseVision is an academic screening and decision-support prototype, not a diagnostic tool.

Table T14 -- CirCor Clinical Outcome Results

Experiment: EXP-C2

Objective: O6 (large-sample validation), O1 (external corpus)

Source: outputs/08_circor_external_validation/T14_circor_outcome_results.csv; outputs/08_circor_external_validation/EXP-C2/per_fold_by_level.csv

Generated by PV-MEPCG / PulseVision at 2026-09-10T16:53:57.586979+00:00